## Blad1
Product Comparison Chart
Manufacturer | ZOLL | Welch Allyn | Philips | Philips | Cardiac Science | Physio-Control | Schiller | HeartSine | NEO TOP | Wienmann | CardiAid | CU Medical
Model | AED Plus | AED 10 | HeartStart FR2 | HeartStart FR3 | Powerheart G3 | CR Plus | FRED easy | Samaritan PAD 500P | HeartPlus | MEDUCORE Easy | CardiAid CT0207 AED | iPAD NF 1200
Assessment & Therapy Comparison
Operation | Semi-automatic | Semi-automatic | Semi-automatic | Semi-automatic | Semi-automatic and manual | Semi - automatic | Semi - automatic | Semi-Automated | automated | automated (one-button-operation) | automated (one-button-operation) | Semi-Automated
Waveform Type | Low energy Rectilinear Biphasic | Biphasic Truncated Exponential (BTE) | Biphasic truncated exponential (BTE) | Biphasic truncated exponential (BTE) | Biphasic Truncated Exponential (BTE) | Biphasic Truncated Exponential (BTE) with voltage and duration compensation for patient impedance | Biphasic pulsed - MULTIPULSE BIOWAVE (patented) | SCOPE (self-compensating output pulse envelope) Biphasic Escalating Waveform. | Biphasic Truncated Exponential (BTE) | Biphasic, current-limited | Biphasic, current-limited | e-cube biphasic; truncated exponential type, impedance compensated
Waveform slope change | No, Superior Low Energy Constant Current Rectilinear Biphasic | Yes, 150 ,200, 300J or 200, 300, 360J user configurable | Yes, variable duration only with the goal of constant energy | N/A | Escalating Variable Energy (VE) | Yes, variable duration only with the goal of constant energy | N/A | Yes, variable duration only with the goal of constant energy | N/A | N/A | Cosntant or escalating , Programmable (factory setting) | N/A
Energy Output | RBW – Fixed or escalating low energy | Variable – User Configurable | 150J, 150J, 150J, minimum 130J | N/A | Variable Energy levels 105J to 360J. 200(VE), 300(VE), (VE) 300 from factory - 5 energy protocols available | User configurable from200J to 350 joules (150J minimum outside the US) | Configurable: Adult: 1 to 150 Joules - Pediatric: 1 to 70 Joules | Adult: 1. Shock 150J - 2. Shock 150J - 3. Shock 200J; Pediatric : 1. Shock 50J - 2. Shock 50J - 3. Shock 50J; | Adult: 180Joules into a 50 Ohm load - children: 50 joules into a 50 Ohm load. | can be pre-configured; low energy (max. 200J) or high-energy (max. 310J), impedance compensated | Low energy (max 181J) or high energy (max 237 J) programmable, energy adjusted to patient impedance | 200 Joules nominal into a 50 Ohm load
V-Fib rhythm detection | >0.10 mV | N/A | >0.10 mV | N/A | 120-240 | >0.08 mV | N/A | N/A | N/A | exceeds AAMI DF 39 requirements | Yes | Yes
V-Tach rhythm detection | >150 | N/A | Adaptive rate | N/A | 120-240 | >120 | N/A | N/A | N/A | exceeds AAMI DF 39 requirements | Yes | Yes
Operator control of analysis sequence | Always operates on auto analyze | Always operates on Auto analyze | 3 analysis options – no auto analyze, full auto analyze, quit CPR | N/A | Always operates on Auto analyze | Always operates on auto analyze | N/A | N/A | N/A | N/A | N/A | N/A
Monitors heart rhythms during CPR | Background monitoring during CPR | N/A | Yes | N/A | Yes (Optional) | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Motion detection | No, noise detection | No | Fixed On, Alerted when ECG analysis is interrupted | N/A | Yes | Optional motion detection during analysis | N/A | N/A | N/A | constant check of signal quality; acoustic warning at patient movement | checks the signal quality; acoustic warning at patient movement | N/A
Asystole detection | Visible w/ECG display | N/A | Not part of algorithm Visible if ECG display is activated. | N/A | 0.08 mV | N/A | N/A | N/A | N/A | < 0.08 mV | Yes | No
CPR Time - countdown timer during which CPR is to be performed | User definable after no shock advised analysis | N/A | User definable (0-180sec 30sec intervals) 2 settings, after shock analysis. Option to quit CPR. | N/A | No | User definable (0-180 seconds) Two settings, after 3 shocks or No Shock advised analysis | N/A | Fixed at 2 minutes. | N/A | N/A | N/A | N/A
Ease of Operation / Ease of Maintenance
Steps to first shock | 1. Remove cover. 2. Push On/Off button.1 3. Apply electrode pads. 4. Press shock button when indicated. | 1. Push On/Off 2. Apply pads to patient. 3. Press shock button when indicated | 1. Push On/Off button. 2. Apply electrode pads to patient. 3. Plug electrode pads into device. 4. Press shock button when indicated. | N/A | 1. Open lid to activate. 2. Apply electrode pads. 3. Press shock button when indicated. | 1. Open Lid to activate, 2. Apply electrode pads to patient, 3. Press shock button when indicated | 1. Switch the unit on, 2. Place the electrodes on the patient's chest as indicated on the pouch, 3. Follow the device's voice and displayed instructions | 3 | N/A | 1. switches on automatically when cover is opened flashing shock key (one-button-operation) Info key | 1. automatic switch on when cover is opened - 2. Flashing shock key (one-button operation) - 3. info Key | N/A
Buttons to Operate | 2 | 4 (2 for configuring) | N/A | N/A | 1 | 2 soft keys | N/A | 2 |  | 1 | 1 | N/A
Voice prompts | Yes, adjustable via configuration. | Yes | Yes, adjustable volume, 2 sets of prompts. Yes | Yes | Yes | Yes, Adjustable volume | Yes | Yes | Yes | Yes | Yes | Yes
Visual prompts | Yes with chain of survival images/ rescuer sequence of action. | Yes | Yes | Yes | Yes | Yes | N/A | Yes No display screen with text Icons with lighted LEDs | Yes, with images | Symbols | Illuminated Symbols, traffic light principle | No
Display # of shocks & energy of last shock | # of Shocks | N/A | Yes | N/A | Yes | No | N/A | No | N/A | N/A | Announcement of the elapsed time and number of shocks since device started, when info-button is pressed | N/A
ELECTRODES
Pre-connected electrodes | Yes | Yes | No | N/A | Yes (non-polarized) | Yes | No | Yes | N/A | N/A | Yes | N/A
One piece electrode | Yes | No | Patent Pending | N/A | No | No | No | No | N/A | N/A | N/A | N/A
Pre-connected electrode quality self-check | Yes | Yes | No | N/A | Yes | No | No | No | N/A | N/A | N/A | N/A
Electrode shelf life | 5 Years | 2 years | 1 year | N/A | 2 years | 2,5 years | 2 years | Samaritan disposable pads - 3.5 years | 2 years | 30 months from date of manufacture - disposable self-adhesive electrodes ready for use, sealed and packed with connector outside of package | 30 months from date to manufacture | 2 years
Device Self-checks | Every 7 days (default) or every 6-5-4-3-2-1 (reconfigurable with ZAS) with full 200J self-test every month | N/A | Fuel Gauge on FR2 Display | N/A | Daily, Weekly, Monthly, Power On | Weekly, power on | N/A | Weekly at midnight on Sunday | N/A | Automatic daily, monthly and when device is switched on | Automatic daily, monthly and when device is switched on | Daily self test
Pediatric Capability
Levels of energy sent to pads | Adult: 120J – 150J – 200J - Pediatric: 50 J - 70J - 85J | Adult: 200J – 300J – 360J - Pediatric: 200 J - 300J - 360J | Adult: 150J – 150J – 150J - Pediatric: 150 J - 150J -150J | N/A | Adult: 200J – 300J – 300J - Pediatric: 200 J - 300J - 300J | Adult: 200J - 300J - 360 J ; Pediatric: 200J - 300J, 360J | Configurable: Adult: 1 to 150 Joules - Pediatric: 1 to 70 Joules | N/A | N/A | N/A | N/A | N/A
Independent adult & pediatric energy protocols | Yes | No | Reduces defibrillator shock energy to nominal 50 Joules into a 50 Ohm load | N/A | No | No | automatic detection of pediatric electrodes | N/A | N/A | N/A | N/A | N/A
Pediatric pad pre-connection | Yes | No | No | No | No | No - Connect during rescue | No | N/A | N/A | N/A | N/A | N/A
Pediatric Electrocardiogram (ECG) Analysis | Yes | No - same as adult | No | N/A | No - same as adult | No - Same as Adult | No | N/A | N/A | N/A | N/A | N/A
Pad-type (adult/pediatric) recognized & announced to prevent errors | Yes | No (voice prompts describe application of adult pads during pediatric rescue with completely different pads) | N/A | N/A | No | No (voice prompts describe application of adult pads during pediatric rescue with completely different pads) | N/A | N/A | N/A | N/A | N/A | N/A
Pediatric defibrillation technology | Configurable pediatric energy only sent to pads | Adult energy sent to pads with hardware in wire to shunt energy | N/A | N/A | Adult energy sent to pads. Hardware in wire shunts energy. | Adult energy sent to pads with hardware in wire to shunt energy | N/A | N/A | N/A | N/A | N/A | N/A
POWER / Battery
Rescue battery indicator | Yes | N/A | Yes | N/A | Yes | Primary Battery (nonrechargeable battery with status indicator) | Yes | N/A | N/A | N/A | N/A | State LED, i-Button and voice prompt
Low battery indicator | Visual & Audio | Yes - Visual | Visual & Audio | N/A | Visual & Audio | Visual | Yes | Visual only | N/A | N/A | N/A | State LED, i-Button and voice prompt
Battery capacity gauge | No | Yes (10 capacity levels) | Fuel Gauge on FR2 Display | N/A | Yes, Smart Gauge | No | N/A | N/A | N/A | N/A | N/A | N/A
Battery type | 10 Lithium Manganese 123A batteries available at retail. | Non-rechargeable Lithium | Lithium Manganese | 12 VDC, Lithium Manganese dioxide. Disposable , long-life primary cell. | IntelliSense Lithium | Lithium Sulfur Chloride (Charge-pak) | Li-MnO2 battery or NiCd rechargeable battery | Lithium Manganese Dioxide 18V, 0.8 Amps Hrs - Battery life: 3 yrs | Non charge type Lithium primary cell, 21Volt DC 1.3 Ah - charge for energy | LiSO2 and Li-Ion | Alkaline | 12 Volt DC, 4.2 Ah, Lithium Manganese Dioxide disposable
Battery Capacity | 5 years (300 shocks or 1.5 hours of continuous monitoring/defibrillation | 5 years (90 discharges at 360J or 125 discharges at 200J or 150 discharges at 150J or 5 hours ECG monitoring) | 4 Years minimum when installed by date, 5 years typical | when new, a minimum of 300 shocks or 12 hours of operating times at 25C. | 4 year (300 shocks or 9 shocks [360J]@ 1% battery life) | 2,5 years, Typical: 30 full discharge or 3.5 hrs of operation/min. 20 full discharges or 2 hrs and 20 min. of operation. | N/A | 3.5 years (number of shocks and continuous monitoring unknown) | N/A | LiSO2 = typically 200 shocks* min. 100 shocks* Li-Ion = typically 100 per charge | typically 210 shocks , min. 100 shocks. | Minimum 200 shocks (10 hours operating time in AED mode)
Charge Time | Less than 10 seconds | N/A | Less than 10 seconds | N/A | 55 seconds for three shock series | 9 to 10 seconds | N/A | Less than 7 seconds (100J) and less than 12 seconds (200J) from a SA message with new Pad-Pak | N/A | typically < 10 seconds | typically < 10 seconds, maximum < 15 seconds | < 10 seconds
Data Management
Fixed internal memory | Yes | Yes | Yes | N/A | Yes (20 minutes ECG and event recording) | Yes | Yes | Yes | Yes | automatic recording of ECG and event data in internal memory | Yes | N/A
Internal memory capacity | 20 minutes, configured w/audio or w/out audio 3.5 hrs | 1MB: 250 4-sec ECG samples or 3000 time stamped events | 1 event. Elapsed time and # of shocks only, but played back on display only – no data download capacity | N/A | 34 minutes ECG data with event annotation | 20 minutes of ECG for the current patient, summarized data stored for the previous patient | Internal memory for the recording of ECG, events and ambient noise | 45 minutes of ECG (full disclosure) and event/incident recording | 20 minutes for ECG, 15 minutes for ambient noise, 500 cases for result of self diagnosis | Up to four data records of up to a max. of 2 hours of complete ECG and event data in internal memory | Automatic recording of ECG and event data in internal memory up to 4 sessions with a total maximum duration of 4 hours | first 40 minutes of ECG and the entire incident's events and analyses decisions
Removable memory | No | No | PC Data Cards - 8 hrs for event data or 1 hr with audio enabled | Data Card SD format. Holds a minimum of 120 minutes of data | No | No | SD card | No | Built in memory - external SD memory support | No - via infrared interface | No | No
Data Transmission | Data is transferred from AED via IrDA port | Data is transmitted via IR Port to PC or Printer with SmartLink software | N/A | Bluetooth wireless transmissions of event data to a Bluetooth compatible device with FTP capability | RescueLink review software sent via Serial or USB port to PC | Serial IrDA link Wireless transfer to PC | Data transmission to a server via a network using the automatic FRED easy Ethernet | Custom USB cable directly connected to PC and SAVER windows-based data review software | N/A | via wireless connection to PC software Easy- View (part of supply schedule) | Bluetooth connection to PC Software CardiAid (only for service) - | Wireless infrared transmission of event data to PC through IrDA port
Ability to store multiple events | No | Yes | Yes, with optional data card. | N/A | Yes | Yes, up to 2 patients can be stored internally | N/A | N/A | N/A | N/A | N/A | N/A
Ability to combine AED data with data from ALS device | Yes, with ZDC or ZDR. | N/A | Yes, compatible with Codemaster data. | N/A | No | Yes, with CodeStat Suite data program | N/A | N/A | N/A | N/A | N/A | N/A
GENERAL
CPR support | Real CPR Help® - Adaptive metronome for 100 bpm. Real-time feedback for compression depth. Voice & text prompts for “Push Harder” and “Good Compressions.” | Timed minute of silence with second to go countdown at LCD screen display | Timed minute of silence (default) or fixed metronome for 100 bpm with voice prompts for breathing at 15-2 (press "I" button) | Timed minute of silence (default) or fixed metronome for 100 bpm with voice prompts for breathing at 15-2 (press "I" button) | Timed minute of silence (default) or " Continue CPR" prompt each 15 seconds (optional) | Timed minute of silence (default) or " Continue CPR" prompt each 15 seconds (optional) | No | Yes | N/A | Yes | N/A | No
Communication with rescuer | Voice prompts, text & circle of picture graphics | Voice prompts, LCD screen text only and 1 picture graphic | Voice prompt and 1 picture graphic | Voice prompt and 1 picture graphic | Voice prompts and 38 character text | Voice prompts and 1 picture graphic | Voice prompts | Voice prompts | Detailed voice messages for Electric Shock and for CPR | Voice Prompt | N/A | Voice Prompt
CPR Assists | Yes | No | No | CPR meter with Q-CPR Technology | No | No | No | Constant beep at 100 compressions per minute with no break for breaths. Re-configuration not possible. | Voice messages | Yes | N/A | No
Airway Support | Yes | No | No | No | No | No | No | No | N/A | N/A | N/A | N/A
Manual Override | No | No | Option | N/A | No | Yes, option | Yes | No | N/A | N/A | N/A | N/A
LCD Text Display | Option | Yes | Standard | N/A | Yes (38 Characters) | Text and graphics | Yes | NO - Icons only | No | N/A | N/A | N/A
Display with ECG | Option | No | Option | N/A | No | Yes, provides nondiagnostic ECG display of the patient's heart rhythm | Yes | No | No | N/A | No | Yes
Device Height | 13.3 cm | 7 cm | 6.6 cm | 7 cm | 8 cm | 10.7 cm | 7 cm | 4.8 cm | N/A | 9.3 cm | 30.4 cm (including handle) | 7 cm
Device Width | 24.1 cm | 17.5 cm | 21.8 cm | 13.5 cm | 27 cm | 20.3 cm | 23 cm | 18.4 cm | N/A | 24 cm | 30.1 cm | 22 cm
Device Depth | 29.2 cm | 21 cm | 21.8 cm | 21.8 cm | 31 cm | 24.1 cm includes handle | 22 cm | 20 cm | N/A | 24 cm (including handle) | 11.2 cm | 26 cm
Size in Cubic Inches | 418 | N/A | 177 | N/A | 434 | 319 | N/A | N/A | N/A | N/A | N/A | N/A
Weight w/battery | 2.37 kg to 2.93 kg | 1.6 kg with battery | 2.1 kg with battery - 1.8 kg without battery | 1.42 kg Defibrillator only, 1.7kg with FR3 standard battery installed | 3.10 kg | 2.04 kg | 1.5 kg with battery | 1.1 kg including Pad-Pak Battery | N/A | 2.6 kg - including battery and electrodes | 3.1 kg including battery and electrodes | 2.2 kg
Operating Temperature | 0 – 50ºC 32 –122 F | 0 – 50ºC 32 –122 F | 0 – 50ºC - standby: 0 - 43ºC recommended storage | N/A | 0 – 43ºC - 32 –109º F | 0 – 50ºC | N/A | 0 – 50ºC | N/A | 0 – 50ºC | 0 – 50ºC | 0 – 40ºC
Ingress Protection (IP) Rating (Dust/Moisture) | IP55 | IP24 | IP55 | IP55 | IP24 | IPX4 | N/A | IP 56 | N/A | IP 54 | IP 4 | IP 54
Carrying Case | Soft Case or wall mount Included | Soft Case Included | Soft Case/ Hard Case | Hard Case included | Soft Case | N/A | N/A | Soft Case included | N/A | N/A | Yes | Option
Warranty | 7 years | 5 years | 5 years | 5 years | 7 years | 8 years | 5 years | 7 Years | N/A | 5 years | 6 years | 5 years